## size_correction
USER INPUTS: | REF TAG SHOULD MATCH CONSTANTS.CSV |  |  |  | COPY AND PASTE IN RAW DATA | [14 empty] | INPUT VALUES FOR DIRECT N2O REF. INJECTION | [5 empty] | corrected to m/z=44 peak area of 20 Vs |  |  | scale normalized 45R and 46R |  | PYTHON/MATLAB INPUTS | [6 empty] | MATLAB/PYTHON OUTPUT |  | SIZE CORRECTED DELTA VALUES |  |  | peak area 20 | N2O CONCENTRATIONS
run_date | ref_tag | Row | Identifier 1 | Is Ref _ | d 15N/14N | d 18O/16O | d 17O/16O | Area 44 | Area 30 | BGD 44 | Rt | FileHeader: Filename | Time Code | rR 45N2O/44N2O sam | rR 46N2O/44N2O sam | rR 31NO/30NO sam | rR 45N2O/44N2O std | rR 46N2O/44N2O std | rR 31NO/30NO std | 31R | 45R | 46R | raw 45rR/45rR | raw 46rR/46rR | raw 31rR/31rR | size corrected 31rR/31rR | size corrected 45rR/45rR | size corrected 46rR/46rR | scale decompressed 45rR/45rR | scale decompressed 46rR/46rR | size corrected 31R | size corrected 45R | size corrected 46R | gamma | kappa | DATE | SAMPLE ID | d15Na | d15Nb | SP | d15Nbulk | d17O | d18O | Pre_weight_g | Post_weight_g | volume analyzed (g) | matrix | volume analyzed (L) | N2O (nmols) | +/- | N2O (nmols/L) | +/-
201204 | ATM | 2 | ATM_EQ_SW | 0 | 5.406 | 46.78 | 3.478 | 4.769 | 4.95 | 2.1 | 1547.6 | 201204_V-0002.dxf | 44169.7648 | 0.76604 | 0.74142 | 0.43208 | 0.76203 | 0.73643 | 0.42841 | 0.0037338 | 0.007741 | 0.0021013 | 1.0053 | 1.0068 | 1.0086 | 1.0094 | 1.0057 | 1.0075 | 1.006 | 1.0071 | 0.003769 | 0.0077876 | 0.0021161 | 0.18326 | 0.09049 | 201204 | ATM_EQ_SW | 15.1055 | -2.7098 | 17.8153 | 6.1978 | 24.0097 | 47.0542 | 265 | 109 | 156 | seawater | 0.15205 | 2.3148 | 0.3276 | 15.224 | 2.1546
201205 | S2 | 1 | S2_ref_1 | 0 | -3.689 | 33.856 | -2.934 | 6.277 | 6.537 | 1.9 | 1546.3 | 201205_V-0000.dxf | 2020/12/05 14:32:14 | 0.75906 | 0.73224 | 0.42948 | 0.76188 | 0.73644 | 0.4293 |  |  |  | 0.9963 | 0.99431 | 1.0004 | 1.0012 | 0.99669 | 0.99498 | 0.9965 | 0.99529 | 0.0037382 | 0.0077139 | 0.0020914 | 0.18326 | 0.09049 | 201205 | S2_ref_1 | 6.8736 | -13.8815 | 20.7551 | -3.5039 | 17.8486 | 34.8799 | 265 | 109 | 156 | DI | 0.15628 | 3.012 | 0.33319 | 19.2729 | 2.132
201205 | B6 | 2 | B6_ref_1 | 0 | 0.364 | 44.33 | 2.266 | 21.185 | 22.028 | 2.0 | 1545.8 | 201205_V-0001.dxf | 2020/12/05 15:07:56 | 0.76223 | 0.73967 | 0.4281 | 0.76191 | 0.73645 | 0.42934 | INPUT SIZE CORRECTION SLOPES |  |  | 1.0004 | 1.0044 | 0.99713 | 0.99706 | 1.0004 | 1.0043 | 1.0004 | 1.004 | 0.0037228 | 0.0077442 | 0.0021098 | 0.18326 | 0.09049 | 201205 | B6_ref_1 | -0.53971 | 1.2825 | -1.8222 | 0.37141 | 22.4614 | 43.9882 | 265 | 109 | 156 | DI | 0.15628 | 9.9049 | 0.38846 | 63.3786 | 2.4856
201205 | ATM | 3 | ATM_EQ_SW_1 | 0 | 6.169 | 46.974 | 3.574 | 4.992 | 5.193 | 1.6 | 1546.9 | 201205_V-0002.dxf | 2020/12/05 15:39:59 | 0.76656 | 0.74167 | 0.433 | 0.762 | 0.73653 | 0.42932 | 31R | 45R | 46R | 1.006 | 1.007 | 1.0086 | 1.0094 | 1.0064 | 1.0077 | 1.0068 | 1.0072 | 0.0037689 | 0.0077935 | 0.0021165 | 0.18326 | 0.09049 | 201205 | ATM_EQ_SW_1 | 14.8516 | -0.85838 | 15.71 | 6.9966 | 24.0954 | 47.224 | 265 | 109 | 156 | seawater | 0.15205 | 2.4179 | 0.32842 | 15.9021 | 2.16
201205 | S2 | 4 | S2_ref_2 | 0 | -2.894 | 35.592 | -2.071 | 36.725 | 38.041 | 2.0 | 1546.2 | 201205_V-0003.dxf | 2020/12/05 16:11:34 | 0.75979 | 0.73348 | 0.42977 | 0.762 | 0.73644 | 0.42892 | 5.6579e-05 | 2.8107e-05 | 4.8711e-05 | 0.9971 | 0.99598 | 1.002 | 1.001 | 0.99663 | 0.99516 | 0.99644 | 0.99546 | 0.0037377 | 0.0077134 | 0.0020918 | 0.18326 | 0.09049 | 201205 | S2_ref_2 | 6.6566 | -13.8026 | 20.4593 | -3.573 | 17.9418 | 35.0634 | 265 | 109 | 156 | DI | 0.15628 | 17.09 | 0.44607 | 109.354 | 2.8543
201205 | B6 | 5 | B6_ref_2 | 0 | -0.866 | 42.325 | 1.272 | 21.034 | 21.866 | 3.5 | 1550.8 | 201205_V-0004.dxf | 2020/12/05 18:31:18 | 0.7614 | 0.73827 | 0.42699 | 0.76201 | 0.73648 | 0.42855 |  |  |  | 0.9992 | 1.0024 | 0.99635 | 0.99629 | 0.99917 | 1.0024 | 0.99912 | 1.0022 | 0.0037199 | 0.0077342 | 0.002106 | 0.18326 | 0.09049 | 201205 | B6_ref_2 | -1.1563 | -0.70441 | -0.45188 | -0.93035 | 21.5158 | 42.1179 | 265 | 109 | 156 | DI | 0.15628 | 9.8351 | 0.3879 | 62.9318 | 2.4821
201205 | ATM | 6 | ATM_EQ_SW_2 | 0 | 4.841 | 44.903 | 2.549 | 4.982 | 5.192 | 2.3 | 1550.5 | 201205_V-0005.dxf | 2020/12/05 19:02:52 | 0.76554 | 0.74014 | 0.43238 | 0.76198 | 0.73648 | 0.4285 |  | SCALE NORMALIZATION 𝜆 |  | 1.0047 | 1.005 | 1.0091 | 1.0099 | 1.0051 | 1.0057 | 1.0054 | 1.0054 | 0.0037708 | 0.0077828 | 0.0021125 | 0.18326 | 0.09049 | 201205 | ATM_EQ_SW_2 | 16.0367 | -4.8657 | 20.9023 | 5.5855 | 23.1173 | 45.2864 | 265 | 109 | 156 | seawater | 0.15205 | 2.4132 | 0.32839 | 15.8717 | 2.1598
201205 |  | 7 | DI | 0 | -9.761 | 44.925 | 2.56 | 17.01 | 17.718 | 2.7 | 1550.8 | 201205_V-0006.dxf | 2020/12/05 19:34:26 | 0.75499 | 0.74002 | 0.43098 | 0.76197 | 0.73649 | 0.4285 |  | 45R | 46R | 0.99084 | 1.0048 | 1.0058 | 1.0059 | 0.99093 | 1.0049 | 0.9904 | 1.0046 | 0.0037559 | 0.0076667 | 0.002111 | 0.18326 | 0.09049 | 201205 | DI | 14.458 | -34.8153 | 49.2733 | -10.1787 | 22.8536 | 44.7645 | 265 | 109 | 156 | DI | 0.15628 | 7.9745 | 0.37298 | 51.0267 | 2.3866
201207 | ATM | 1 | ATM_EQ_SW_1 | 0 | 5.499 | 44.616 | 2.407 | 4.653 | 4.86 | 2.2 | 1550.3 | 201207_V-0000.dxf | 2020/12/07 11:54:27 | 0.76571 | 0.73974 | 0.43346 | 0.76168 | 0.73626 | 0.42919 |  | 1.058 | 0.93784 | 1.0053 | 1.0047 | 1.01 | 1.0108 | 1.0057 | 1.0055 | 1.0061 | 1.0051 | 0.0037742 | 0.0077879 | 0.0021121 | 0.18326 | 0.09049 | 201207 | ATM_EQ_SW_1 | 17.1518 | -4.5745 | 21.7264 | 6.2886 | 22.9885 | 45.0315 | 265 | 109 | 156 | seawater | 0.15205 | 2.2611 | 0.32717 | 14.8712 | 2.1518
201207 | S2 | 2 | S2_ref_1 | 0 | -3.908 | 32.678 | -3.52 | 7.684 | 8.013 | 2.3 | 1549.8 | 201207_V-0001.dxf | 2020/12/07 12:26:40 | 0.75864 | 0.73124 | 0.42947 | 0.76164 | 0.73626 | 0.42925 |  |  |  | 0.99607 | 0.99318 | 1.0005 | 1.0012 | 0.99641 | 0.99378 | 0.9962 | 0.99417 | 0.0037382 | 0.0077116 | 0.002089 | 0.18326 | 0.09049 | 201207 | S2_ref_1 | 7.0099 | -14.5751 | 21.585 | -3.7826 | 17.2545 | 33.7094 | 265 | 109 | 156 | DI | 0.15628 | 3.6625 | 0.3384 | 23.4356 | 2.1654
201207 | B6 | 3 | B6_ref_1 | 0 | 0.138 | 44.011 | 2.108 | 20.544 | 21.38 | 2.4 | 1548.9 | 201207_V-0002.dxf | 2020/12/07 13:08:33 | 0.76178 | 0.73897 | 0.42797 | 0.76164 | 0.73598 | 0.42905 |  |  |  | 1.0002 | 1.0041 | 0.99749 | 0.99746 | 1.0002 | 1.004 | 1.0002 | 1.0038 | 0.0037243 | 0.0077425 | 0.0021092 | 0.18326 | 0.09049 | 201207 | B6_ref_1 | 0.082242 | 0.22102 | -0.13878 | 0.15163 | 22.3259 | 43.7202 | 265 | 109 | 156 | DI | 0.15628 | 9.6085 | 0.38608 | 61.4822 | 2.4704
201207 | ATM | 13 | ATM_EQ_SW_2 | 0 | 4.622 | 44.532 | 2.365 | 4.851 | 5.042 | 2.2 | 1554.3 | 201207_V-0012.dxf | 2020/12/07 19:56:02 | 0.76494 | 0.73974 | 0.43109 | 0.76155 | 0.73634 | 0.42775 |  |  |  | 1.0045 | 1.0046 | 1.0078 | 1.0087 | 1.0049 | 1.0054 | 1.0052 | 1.005 | 0.0037662 | 0.007781 | 0.0021119 | 0.18326 | 0.09049 | 201207 | ATM_EQ_SW_2 | 14.3938 | -3.6809 | 18.0747 | 5.3564 | 22.9446 | 44.9445 | 265 | 109 | 156 | seawater | 0.15205 | 2.3527 | 0.3279 | 15.4733 | 2.1566
201207 | ATM | 23 | ATM_EQ_SW_3 | 0 | 5.383 | 45.667 | 2.928 | 4.602 | 4.783 | 1.8 | 1551.9 | 201207_V-0022.dxf | 2020/12/08 01:14:46 | 0.7656 | 0.74048 | 0.43216 | 0.76163 | 0.73627 | 0.42803 |  |  |  | 1.0052 | 1.0057 | 1.0096 | 1.0105 | 1.0056 | 1.0065 | 1.006 | 1.0061 | 0.003773 | 0.0077872 | 0.002114 | 0.18326 | 0.09049 | 201207 | ATM_EQ_SW_3 | 16.6798 | -4.3336 | 21.0135 | 6.1731 | 23.4877 | 46.0199 | 265 | 109 | 156 | seawater | 0.15205 | 2.2375 | 0.32698 | 14.7162 | 2.1505
201207 | S2 | 24 | S2_ref_2 | 0 | -4.002 | 33.719 | -3.002 | 8.283 | 8.603 | 2.0 | 1551 | 201207_V-0023.dxf | 2020/12/08 01:46:13 | 0.75863 | 0.73199 | 0.42839 | 0.76168 | 0.73628 | 0.42812 |  |  |  | 0.996 | 0.99417 | 1.0006 | 1.0013 | 0.99633 | 0.99474 | 0.99612 | 0.99507 | 0.0037386 | 0.007711 | 0.0020909 | 0.18326 | 0.09049 | 201207 | S2_ref_2 | 7.103 | -14.8931 | 21.996 | -3.895 | 17.7355 | 34.6569 | 265 | 109 | 156 | DI | 0.15628 | 3.9395 | 0.34063 | 25.2077 | 2.1796
201207 |  | 25 | DI | 0 | -7.13 | 44.149 | 2.176 | 13.6 | 14.154 | 2.0 | 1551.2 | 201207_V-0024.dxf | 2020/12/08 02:17:40 | 0.7566 | 0.73928 | 0.43062 | 0.7617 | 0.73627 | 0.42825 |  |  |  | 0.99332 | 1.0041 | 1.0055 | 1.0059 | 0.99349 | 1.0044 | 0.99312 | 1.0041 | 0.0037558 | 0.0076878 | 0.00211 | 0.18326 | 0.09049 | 201207 | DI | 13.7252 | -28.3354 | 42.0606 | -7.3051 | 22.5651 | 44.1934 | 265 | 109 | 156 | DI | 0.15628 | 6.3979 | 0.36034 | 40.9382 | 2.3057

## scale_normalization
USER INPUT: CALIBRATED DELTA VALS FOR EACH REFERENCE MATERIAL | [7 empty] | CALCULATE KNOWN 45R/45R & 46R/46R FOR EACH REFERENCE MATERIAL
Calibrated values | d15Na | d15Nb | d15Nbulk | d18O | SP | 15Ralpha | 15Rbeta | 18R | 17R | 45R | 46R | 45R/45R | 46R/46R
ATM | 15.7 | -3.3 | 6.2 | 44.3 | 19 | 0.0037342 | 0.0036644 | 0.002094 | 0.00038849 | 0.0077871 | 0.0021106 | 1.0059 | 1.0044
S2 | 5.55 | -12.87 | -3.66 | 32.73 | 18.42 | 0.0036969 | 0.0036292 | 0.0020708 | 0.00038627 | 0.0077124 | 0.0020871 | 0.9963 | 0.99323
B6 | -0.40397 | -0.14815 | -0.27606 | 41.9493 | -0.25582 | 0.003675 | 0.003676 | 0.0020893 | 0.00038804 | 0.007739 | 0.0021057 | 0.99974 | 1.0021
CA06261 (EMPA STD1, "Joni") | -22.21 | -49.28 | -35.745 | 26.94 | 27.07 | 0.0035948 | 0.0034953 | 0.0020592 | 0.00038515 | 0.0074753 | 0.0020745 | 0.96567 | 0.98726
53504 (EMPA STD2) | 1.71 | 94.44 | 48.075 | 36.01 | -92.73 | 0.0036828 | 0.0040237 | 0.0020774 | 0.0003869 | 0.0080934 | 0.0020952 | 1.0455 | 0.9971
CA08214 (EMPA STD3) | 17.11 | -3.43 | 6.84 | 35.29 | 20.54 | 0.0037394 | 0.0036639 | 0.002076 | 0.00038676 | 0.0077901 | 0.0020925 | 1.0063 | 0.99583
Log. normalization following Kaiser et al. (2007) Triple oxygen isotope analysis of nitrate using the denitrifier method and thermal decomposition of N2O. Anal. Chem. 79, 599-607 doi:10.1021/ac061022s
USER INPUT: REFERENCE APPROPRIATE VALUES IN COLUMNS M & N
Identifier 1 | known 45rR/45rR | known 46rR/46rR | size corrected 45rR/45rR | size corrected 46rR/46rR | ln(size corrected 45rR/45rR) | ln(size corrected 46rR/46rR) | ln(known 45rR/45rR) | ln(known 46rR/46rR) | 45R 𝜆 | 46R 𝜆
ATM_EQ_SW | 1.0059 | 1.0044 | 1.0057 | 1.0075 | 0.005669 | 0.0074937 | 0.005932 | 0.0044128 | 1.058 | 0.93784
S2_ref_1 | 0.9963 | 0.99323 | 0.99669 | 0.99498 | -0.0033158 | -0.0050357 | -0.0037106 | -0.0067895
B6_ref_1 | 0.99974 | 1.0021 | 1.0004 | 1.0043 | 0.0003832 | 0.0043014 | -0.00026012 | 0.0020837
ATM_EQ_SW_1 | 1.0059 | 1.0044 | 1.0064 | 1.0077 | 0.0063856 | 0.0076776 | 0.005932 | 0.0044128
S2_ref_2 | 0.9963 | 0.99323 | 0.99663 | 0.99516 | -0.0033737 | -0.0048489 | -0.0037106 | -0.0067895
B6_ref_2 | 0.99974 | 1.0021 | 0.99917 | 1.0024 | -0.00082966 | 0.0023832 | -0.00026012 | 0.0020837
ATM_EQ_SW_2 | 1.0059 | 1.0044 | 1.0051 | 1.0057 | 0.0050812 | 0.0056954 | 0.005932 | 0.0044128
DI |  |  | 0.99093 | 1.0049 | -0.0091142 | 0.0049257
ATM_EQ_SW_1 | 1.0059 | 1.0044 | 1.0057 | 1.0055 | 0.005703 | 0.0054477 | 0.005932 | 0.0044128
S2_ref_1 | 0.9963 | 0.99323 | 0.99641 | 0.99378 | -0.0035946 | -0.0062381 | -0.0037106 | -0.0067895
B6_ref_1 | 0.99974 | 1.0021 | 1.0002 | 1.004 | 0.00017967 | 0.0040262 | -0.00026012 | 0.0020837
ATM_EQ_SW_2 | 1.0059 | 1.0044 | 1.0049 | 1.0054 | 0.0048687 | 0.0053457 | 0.005932 | 0.0044128
ATM_EQ_SW_3 | 1.0059 | 1.0044 | 1.0056 | 1.0065 | 0.0056229 | 0.0064468 | 0.005932 | 0.0044128
S2_ref_2 | 0.9963 | 0.99323 | 0.99633 | 0.99474 | -0.0036735 | -0.0052698 | -0.0037106 | -0.0067895
DI |  |  | 0.99349 | 1.0044 | -0.0065263 | 0.0043905

## scrambling
Template for algebraic solution for gamma and kappa. | [11 empty] | 17RVSMOW | 0.0003799 | =LAMBDA(R18,n,R45,R46,D17O,k,R18VSMOW,R17VSMOW, IF(n>0,Rsolve18(LET(R17,(1 + D17O / 1000) * RVSMOW17 * (Rsam18/RVSMOW18) ^ k, R46 - (R45 - R17) * R17 - (R45 - R17) ^ 2/4),n-1,R45,R46,D17O,k,R18VSMOW,R17VSMOW), R18))
Both algebraic and least squares methods are also available in the "Scrambling" function of pyisotopomer. | [11 empty] | 18RVSMOW | 0.0020052
Copy and paste in values for ref #1: atmosphere-equilibrated seawater | [5 empty] | Copy and paste in values for ref #2: B6 reference gas | [5 empty] | k | 0.516
15Ralpha, ref1 | 15Rbeta, ref1 | 31Rm, ref1 | 45Rm, ref1 | 46Rm, ref1 | D17O_ref1 | 15Ralpha, ref2 | 15Rbeta, ref2 | 31Rm, ref2 | 45Rm, ref2 | 46Rm, ref2 | D17O_ref2 | R18_ref1 | R18_ref2 | 17R, ref1 | 17R, ref2 | gamma | kappa
0.0037342 | 0.0036644 | 0.0037689 | 0.0077935 | 0.0021165 | 0 | 0.003675 | 0.003676 | 0.0037228 | 0.0077442 | 0.0021098 | 0 | 0.0020999 | 0.0020934 | 0.00038905 | 0.00038843 | 0.20872 | 0.11597
0.0037342 | 0.0036644 | 0.0037689 | 0.0077935 | 0.0021165 | 0 | 0.003675 | 0.003676 | 0.0037199 | 0.0077342 | 0.002106 | 0 | 0.0020999 | 0.0020897 | 0.00038905 | 0.00038807 | 0.1733 | 0.079868
0.0037342 | 0.0036644 | 0.0037708 | 0.0077828 | 0.0021125 | 0 | 0.003675 | 0.003676 | 0.0037228 | 0.0077442 | 0.0021098 | 0 | 0.002096 | 0.0020934 | 0.00038868 | 0.00038843 | 0.17734 | 0.08459
0.0037342 | 0.0036644 | 0.0037708 | 0.0077828 | 0.0021125 | 0 | 0.003675 | 0.003676 | 0.0037199 | 0.0077342 | 0.002106 | 0 | 0.002096 | 0.0020897 | 0.00038868 | 0.00038807 | 0.14191 | 0.04849
0.0037342 | 0.0036644 | 0.0037742 | 0.0077879 | 0.0021121 | 0 | 0.003675 | 0.003676 | 0.0037243 | 0.0077425 | 0.0021092 | 0 | 0.0020955 | 0.0020929 | 0.00038863 | 0.00038838 | 0.14995 | 0.05762
0.0037342 | 0.0036644 | 0.0037662 | 0.007781 | 0.0021119 | 0 | 0.003675 | 0.003676 | 0.0037243 | 0.0077425 | 0.0021092 | 0 | 0.0020953 | 0.0020929 | 0.00038862 | 0.00038838 | 0.26277 | 0.17042
0.0037342 | 0.0036644 | 0.003773 | 0.0077872 | 0.002114 | 0 | 0.003675 | 0.003676 | 0.0037243 | 0.0077425 | 0.0021092 | 0 | 0.0020975 | 0.0020929 | 0.00038882 | 0.00038838 | 0.16881 | 0.076477
[16 empty] | 0.18326 | 0.09049
[16 empty] | 0.041096 | 0.041284

## concentration_constants
CONSTANTS USED IN CONCENTRATION CALCULATIONS
nmol per Vs | 0.46236 | *x + | 0.10976 |  | Equivalent to: | 2.1628 | Vs/nmol
error | 0.0080187 |  | 0.30904
seawater kg/L | 1.026 |  |  | DI water kg/L | 0.9982 |  | freshwater kg/L | 1.002
seawater L/kg | 0.97466 |  |  | DI water L/kg | 1.0018 |  | freshwater L/kg | 0.998
seawater L/g | 0.00097466 |  |  | DI water L/g | 0.0010018 |  | freshwater L/g | 0.000998
Linest results
| SLOPE | INTERCEPT
| 0.46236 | 0.10976
ERROR | 0.0080187 | 0.30904
| 0.99343 | 0.96347
| 3324.731 | 22
| 3086.2689 | 20.4221